## Slide 1
Developed for the
Department of Environmental Protection
Bureau of Radiation Protection
Refresher Training for
X-Ray Equipment Operators
Presented by
David R. Simpson, CHP, PhD
Bloomsburg University

## Slide 2: Outline of Course
Part 1: Introduction and Overview of the Training Program
Part 2: Fundamentals of Radiation Science
Part 3: Review of X-ray Imaging
Part 4: Safety Plan, Documentation and QA
Part 5: Regulations

## Slide 3: Introduction and Overview of Training Program
Part 1

## Slide 4: Introduction and Background
X-rays a key component of diagnostic medicine for many years
BUT will result in exposure of patient and staff to ionizing radiation
Food and Drug Administration (FDA) recommends two steps for medical X-rays:
  Justification of procedure
  Optimization of procedure

## Slide 5: Justification of Imaging Procedure
Imaging procedure should do more good than harm; therefore, exams should be performed only when necessary to:
  Answer a medical question
  Treat a disease
  Guide a procedure

## Slide 6: Optimization of Imaging Procedure
Exams should use techniques that are adjusted to:
  Administer lowest radiation dose that yields image quality adequate for diagnosis or intervention
That is, radiation doses should be “As Low As Reasonably Achievable,” referred to as the ALARA Principle

## Slide 7: Purpose of Training
According to the National Council on Radiation Protection and Measurements (NCRP Report No. 134) there are four major reasons for training:
1. Development of worker skills so that tasks may be performed efficiently and with confidence

## Slide 8: Purpose of Training
2. Individuals aware of risk of exposure become active participants in accepting and, where possible, reducing those risks
3. Number and seriousness of accidents can be reduced
4. Workers will be aware of regulatory requirements involved with radiation exposure

## Slide 9: State Requirements for Training
Pennsylvania Department of Environmental Protection (DEP) regulations require individuals operating X-ray equipment to:
  Receive initial instructions in safe operating procedures
  Be competent in the safe use of equipment
  Receive continuing education

## Slide 10: State Requirements for Continuing Education
DEP Technical Guidance Document No. 291-4200-001 lists the following topics for continuing education:
  Basic properties of radiation
  Units of measurement
  Sources of radiation exposure
  Methods of radiation protection
  Biological effects of radiation exposure

## Slide 11: Topics for Continuing Education
  X-ray equipment
  Image recording and processing
  Patient exposure and positioning
  Procedures
  Quality assurance program
  Regulations

## Slide 12: To provide the generic portions of this training for operators performing low-risk medical procedures
Training Goal

## Slide 13: Fundamentals of Radiation Science
Part 2

## Slide 14
Properties of Radiation
Units of Measurement
Sources of Radiation
Biological Effects of Radiation
Fundamental Properties of Radiation for X-ray Imaging

## Slide 15: Ionizing Radiation
High energy particles or electromagnetic (EM) energy
Capable of removing orbital electrons from atoms
Effect called ionization
Resulting atom and electron called ion pair

## Slide 16
Particulate vs. EM Ionizing
Radiation
EM
  Includes gamma, and X-rays
  Generally more difficult to shield
  X-rays are major tool in diagnostic medicine
Particulate
  Includes alpha, beta, neutrons
  Usually easier to shield
  Not used in diagnostic medicine (some applications in therapy)

## Slide 17: Types of EM Radiation
Most EM radiation is non-ionizing
Common names for various energies of EM radiation:
  Radio waves (very low energy, non-ionizing)
  Microwaves (non-ionizing despite common perception)
  Infra-red, visible, and Ultraviolet (non-ionizing, though intense forms may damage skin or eyes)
  X-rays (ionizing)
  Gamma rays and cosmic rays (ionizing)

## Slide 18: EM Radiation Types

## Slide 19: X-rays vs. Other Forms of Ionizing Radiation
X-rays (with some minor exceptions) are produced by machines.
Particulate radiations and gamma rays primarily come from radioactive materials.
X-rays cannot make a person radioactive and cannot result in contamination (loose radioactive materials).

## Slide 20
X-ray machines have an on-off switch
They can be immediately turned off, removing the source of radiation
X-rays vs. Other Forms of Ionizing Radiation

## Slide 21: Quantities & Their Units of Measurement
Three Basic Quantities
  Exposure – measure of charge in air produced by X-ray or gamma radiation
  Absorbed Dose – measure of energy deposited by any type of ionizing radiation in any material
  Equivalent Dose – measure of biological damage to the human caused by various types of ionizing radiation

## Slide 22: Why Three Quantities?
Exposure
  Easy to measure using inexpensive instruments
  Can be related to other two quantities
Absorbed Dose
  Widely applicable to measuring effects of radiation
  Often difficult to measure
Equivalent Dose (or Dose equivalent)
  Allows for biological differences in humans for different types of radiations
  Used for regulatory limits

## Slide 23: Exposure Units
Roentgen
  Oldest unit, defined as 0.000258 coulombs/kg in air
  Still widely used in USA but not defined in the current international scientific unit system (SI)
  Symbol: R
Air kerma (closest SI equivalent)
  Defined as 1 joule/kg in air
  Symbol: Gya (grays in air – see absorbed dose)
To a close approximation 1R = 0.01 Gya

## Slide 24: Absorbed Dose Units
rad
  Traditional unit defined as 100 ergs/g of energy deposited by any type of ionizing radiation in any mass (g) of material
Gray (Gy)
  SI unit defined as 1 joule/kg of energy deposited by any type of ionizing radiation in any mass (kg) of material
1 rad = 0.01 Gy

## Slide 25: Effective Dose Equivalent Dose Units
rem – traditional unit defined as absorbed dose in rad multiplied by modifying factors to account for responses in the human
Sievert (Sv) – SI unit defined as absorbed dose in gray multiplied by modifying factors to account for responses in the human
1 rem = 0.01 Sv

## Slide 26: Relationship of Units
For X-ray radiation, to a close approximation we can assume that:
  1 R = 1 rad = 1 rem
  and
  1 Gya = 1 Gy = 1 Sv

## Slide 27: Common Numerical Prefixes
Numeric Value | Prefix | Symbol
106 | mega- | M
103 | kilo- | k
10-2 | centi- | c
10-3 | milli- | m
10-6 | micro- | m
10-9 | nano- | n

## Slide 28: Dose and Dose Rate
Radiation often measured as a dose rate (dose per time) so dose received is calculated as follows:
Dose received = Dose Rate x Time exposed

## Slide 29: Example of Dose Rate & Conversions
Exposure rate in area is 2 mR/hr.
Is yearly limit of 5 rem exceeded?
(Assume occupancy of 40 hours/week and 50 work weeks in year).
Exposure = 2mR/hr x 40 hr/wk x 50 wks/yr
Exposure = 4000 mR/yr = 4 R/yr
Assume 1R = 1 rad = 1 rem, so
Equivalent dose = 4 rem/yr; limit not exceeded

## Slide 30: Sources of Radiation
Radiation exposure to the general public comes from a number of sources:
  Natural background radiation (in soil, air, our bodies, etc.)
  Medical procedures
  Occupational exposures
  Consumer products
Following slide shows breakdown of exposures and recent trends

## Slide 32: So what has happened?
Background levels are the same, but in % they have decreased to only about half of the total average exposure
Occupational and consumer product levels have remained very low
Medical exposures have increased significantly (on the average)
Why?

## Slide 34: So what has changed?
Large increase in number of CT scans
Increase in nuclear medicine procedures
Newer techniques involve higher doses to the patient

## Slide 35: High-Risk vs. Low-Risk Procedures
Pennsylvania has divided X-ray medical procedures into two risk classes:
  High-risk procedures – utilize energies of < 1 MeV that could exceed skin doses of 200 rads (2 Gy).
    Examples: CT scans, interventional radiography
  Low-risk procedures – any radiologic procedure that is not a “high-risk” procedure.
    Examples: conventional x-rays, dental, podiatry, chiropractic, and veterinary
High-risk results in more dose per patient, but many more patients receive low-risk procedures (especially children and young adults).

## Slide 36: Biological Effects of Radiation
Harmful effects discovered very early
  (Thomas Edison ceased work on X-rays in 1904 following a serious injury and subsequent death of his assistant due to radiation)
“Law of Bergonie and Tribondeau” developed in France in 1906:
  “The more rapidly a cell is dividing, the greater the sensitivity to radiation”
  Not always true but helpful in explaining effects on certain organs such as skin, blood-forming organs, gonads and unborn children.

## Slide 37: What is the primary biological effect?
Research has found that the primary hazard is to DNA
Can result in cell death leading to organ failure, illness, and possible death of exposed person
Or can result in cell mutations leading to cancer in exposed person and possible genetic effects to future generations
Greatest risk at low doses is cancer to exposed person – regulations based on this risk

## Slide 38: Acute vs. Delayed Effects
Acute effects (also referred to as early effects or deterministic effects)
  Typically occur at high doses and appear within days or weeks of exposure
  Examples: skin burns, sterility, loss of hair, etc.
Delayed effects (also referred to as late effects or stochastic effects)
  Occur at low doses over long periods of time
  Cancer is greatest concern

## Slide 39: Goal of Radiation Regulations
Prevent acute effects to exposed person
  Shield sensitive organs such as lens of eye, thyroid, unborn child
Reduce likelihood of delayed effects
  Keep doses low and make sure they are justified
In other words, ALARA:
As low as reasonably achievable

## Slide 40: Review of X-ray Imaging
Part 3

## Slide 41: Basic Components of X-ray Tube
Cathode – heated wire to produce large source of electrons and focusing cup to direct them
Anode – target of high atomic number struck by electrons to produce X-rays
Voltage supply – high voltage supply to accelerate electrons from cathode to anode
Envelope – glass or metal vacuum tube containing anode and cathode
Tube housing – shielding around envelope to protect tube and shield unwanted X-rays

## Slide 43: Other Parts of X-ray Machine
Collimator – restricts X-ray beam to only the area of interest
Filtration – removes unwanted low energy X-rays
Transformer – converts low voltage to high voltage needed for tube
Rectifier – converts AC input voltage to DC needed for X-ray tube

## Slide 44: Primary Settings on X-ray Tube
Current to cathode – measured in milliamps (mA)
Timer for current in seconds
  Current and timer together are measured in mAs
Voltage across the cathode and anode – measured in kVp (kilovoltage peak, the maximum possible energy a photon exiting the X-ray tube can reach)
  kVp determines the energy of electrons which is directly related to energy of X-rays produced

## Slide 45: Quantity and Quality of X-rays
Medical X-rays characterized by quantity and quality
Quantity – number of X-rays reaching the patient
Quality – penetrability or ability of X-ray beam to pass through tissue
(low quality X-rays have little chance of penetrating so they deliver dose to patient while providing no useful medical information)

## Slide 46: Factors Affecting Quantity
mAs | when increased | Increases quantity proportionately
kVp | when increased | Increases quantity by square law
distance | when increased | Decreases quantity by inverse square law
filtration | when increased | Decreases quantity

## Slide 47: Factors Affecting Quality
mAs | when changed | does not change quality
kVp | when increased | improves quality
distance | when changed | does not change quality
filtration | when increased | increases quality

## Slide 48: Impact of Digital Imaging
Digital (computer) imaging replacing conventional (film) systems
  Images immediately available, can be stored and transmitted electronically and post processed to improve image after the fact
Should result in lower dose to patients due to fewer retakes

## Slide 49: Dose Creep in Digital Imaging
In some cases doses increased in digital systems
Effect called dose creep
  Digital resulted in good images without changing factors, even if patient dose was higher, so factors were not optimized to lower dose.
  Also, techniques were used to reduce signal noise that resulted in increased dose.
Manufacturer’s recommendations should be reviewed to minimize dose to patient

## Slide 50: Reducing Patient (and employee) Dose from Medical X-rays
Sources of radiation in Medical X-ray Imaging:
  Primary beam – also called the useful beam; the X-ray beam coming from the tube, through the patient, to the image receptor.
  Scatter radiation – radiation resulting from the primary beam interacting with other materials. The patient is often the largest source of scatter.
  Leakage radiation – leakage radiation from tube housing, generally a minor source from a properly housed tube.

## Slide 51: Administrative Controls for Reducing Radiation Doses
Chapter 221.11 of the Pa. Code, Title 25 - Environmental Protection, has a list of responsibilities for registrants to ensure doses are ALARA
These responsibilities are summarized in the following tables:

## Slide 52
Section Number | Summary of Requirement | Additional Comments and Clarifications
221.11(a) | Registrant is responsible for directing operation and assuring requirements are met. | 
221.11(b) | Operator(s) shall be instructed in safe operating procedures and competent to use equipment. | Instructions shall include items in Appendix A of this section.
221.11(c) | Chart specifying techniques performed with system and technique charts for patient size (ex., adult vs. child) shall be provided in vicinity of control panel. | Chart (or protocol information) should contain pertinent information to particular exams.
221.11(d) | Written safety procedures and rules shall be available- including restrictions for safe use. | Operator shall be able to demonstrate familiarity with these rules.
221.11(e) | Only staff and others required for procedure or training shall be in room during the exposure. | Exception may be made for other patients in room that cannot be moved out (see 221.11(e)(3) below).
221.11(e)(1) | Except for patient, individuals shall be positioned so that no body part will be struck by the useful beam unless protected by 0.5 mm lead equivalent material. | Lead equivalent of material determined at 60kV.

## Slide 53
Section Number | Summary of Requirement | Additional Comments and Clarifications
221.11(e)(2) | Personnel required for exam shall be protected by protective aprons or barriers of at least 0.25 mm lead equivalent or not in direct line of useful beam and at least 2 meters away | Two meter distance is based on nearest portion of body from both tube head and nearest edge of image receptor
221.11(e)(3) | Other patient(s) in room that cannot be moved shall be protected by barriers of at least 0.25 mm lead or equivalent material; positioned out of direct line of the useful beam; and at least 2 meters away. | Again, two meter distance is based on nearest portion of body from both the tube head and nearest edge of image receptor.
221.11(e)(4) | No individual except patient being examined may be in useful beam, unless required to conduct procedure | 
221.11(f) | When patient’s gonads are in useful beam, gonad shielding of at least 0.5 mm lead equivalent shall be used unless it interferes with procedure. | 
221.11(g) | Individuals may not be exposed to useful beam except for healing arts purposes or approved research (see 221.15) | Specifically prohibited are exposures for: training, demonstrations, other non-healing purposes. Exposures for screening purposes must be approved (see 221.13)

## Slide 54
Section Number | Summary of Requirement | Additional Comments and Clarifications
221.11(h) | If patient or image receptor requires auxiliary support during exposure: (1) Mechanical holding devices shall be used when technique permits, (2) Human holder shall be protected per 221.11(e), (3) Individual may not be used to routinely hold image receptors or patients. | 
221.11(i) | Procedures and auxiliary equipment for minimizing patient and personnel exposure commensurate with needed diagnostic information shall be utilized. | 
221.11(j) | Screen and film systems used shall be spectrally compatible. | Defective screens may not be used for diagnostic screening.
221.11(k) | Film may not be used without intensifying screens for routine diagnostic imaging. | An exception to this is for intraoral radiography.
221.11(l) | Shall have a documented QA program in accordance with guidelines established by DEP or by an appropriate organization recognized by DEP. At a minimum, QA program shall address: (1) Repeat rate, (2) Image recording, processing, and viewing, (3) Maintenance and modifications to QA program | Records shall be maintained for inspection by DEP for 3 years. DEP’s guidelines and list of recognized organizations are available on DEP’s website and on request.
221.11(m) | Neither the X-ray tube housing nor collimating device may be hand-held during exposure. | 

## Slide 55
kVp and mAs settings
  Generally lowest dose is achieved when kVp settings are increased and mAs settings decreased within the limits of obtaining a good image.
  Prevent dose creep by “technique creep” – gradually increase kVp while decreasing mAs over successive exams as long as image quality remains satisfactory until optimum reached.
Additional Techniques for
Reducing Dose

## Slide 56
Use of grids
  While grids can reduce scatter radiation and increase image quality, they increase dose to patient.
  Use only as necessary and avoid on children.
Reducing Dose from scatter radiation
  Workers should increase distance from patient, especially from beam side.
  Use carbon fiber or similar material for cassettes, grids and tables.
Additional Techniques for
Reducing Dose

## Slide 57
Collimation and Beam Size
  Primary beam should be sized to cover area of interest but not overly exceed it.
  Collimation to the clinical region of interest should be performed prior to patient exposure.
Additional Techniques for
Reducing Dose

## Slide 58: Special Cases - Children
Historically, children have been imaged with settings similar to adults – resulting in unnecessary dose.
IMAGE GENTLY campaign designed to make medical staff and parents aware of potentially unnecessary exposure to children.
See website at: www.pedrad.org/associations/5364/ig/home

## Slide 62: Additional Sources of Information
International Atomic Energy Agency has PowerPoints available for free:
http://rpop.iaea.org/RPOP/Content/AdditionalResources/Training/1 TrainingMaterial/PaediatricRadiology.htm
Poster from site:

## Slide 63: Other Special Cases – Patient Size
  Larger patients often require adjustments to kVp and mAs that result in higher dose.
  Example: An increase of body thickness from 16 to 24 cm increases scatter 5X.
  Staff should take additional precautions such as increasing distance from patient or use of portable shields.

## Slide 64: Pregnant Patients
Special care should be taken to reduce dose to fetus:
  Females of child-bearing age should be informed of risks.
  Warning signs should be posted.
  If X-ray is necessary, efforts should be made to reduce dose to fetus through shielding and/or positioning.

## Slide 65: Pregnant Workers
Female medical staff required by federal and state regulations to be instructed on risks.
Special limits apply to “declared” pregnant female (woman has option to declare pregnancy).
NRC Regulatory Guide 8.13 provides information on risks, instructions on declaring pregnancy, and lower limits that then apply.

## Slide 66: Radiation Safety Plan, Documentation and QA/QC
Part 4

## Slide 67: Radiation Safety Plan
Mechanism that ensures registrant properly directs X-ray program.

## Slide 68: Guidance for Diagnostic & Interventional X-ray Procedures
Plan should ensure that:
  Radiation activities are performed in accordance with existing laws and regulations.
  Staff are equipped with knowledge of available options regarding risk vs. benefit determinations and appropriate examinations.
  X-ray users and surrounding public receive adequate protection.
From Federal Guidance Report No. 14

## Slide 69: Documentation of Medical X-ray Program
Charts, records, procedures and other documentation are essential for safety and compliance.
Pennsylvania regulations require certain documentation.

## Slide 70: Pennsylvania Requirements for Documentation
25 Pa. Code § 221.11(d)
  Written safety procedures and rules available at facility.
  Operators able to demonstrate familiarity with rules.
  Procedures and rules should be specific for particular applications planned at facility.

## Slide 71: Pennsylvania Requirements for Documentation
25 Pa. Code § 221.11(c)
  Chart (protocol) provided in vicinity of control panel specifying techniques for exams on that system.
25 Pa. Code § 221.12
  Registrant shall maintain records of surveys, calibrations, maintenance, and modifications including names of person performing service.
  Records kept for inspection for minimum of 5 years.

## Slide 72
See 25 Pa. Code, Chapter 221 (X-rays in the Healing Arts) for other documentation requirements for specific applications.
Pennsylvania Requirements for Documentation

## Slide 73: Quality Assurance
According to FDA Regulations, Quality Assurance is defined as:
“…the planned and systematic actions that provide adequate confidence that a diagnostic X-ray facility will produce consistently high-quality images with minimum exposure of the patients and healing arts personnel.”

## Slide 74: FDA Statement on QA Actions
Should include both “quality control” techniques and “quality administration” procedures:
  QC techniques – techniques used in monitoring or testing and maintenance. Concerned directly with equipment.
  QA procedures – management actions intended to guarantee monitoring techniques are properly performed and evaluated and corrective actions taken.

## Slide 75
25 Pa. Code, § 221.11(l) states:
“…the registrant shall have a quality assurance program. This quality assurance program shall be documented and in accordance with guidelines…”
Pennsylvania Regulatory Requirements for QA Program

## Slide 76: Guidelines for Developing QA Program
Available on DEP’s website at: www.dep.state.pa.us/brp/Radiation_Control_Division/X-Ray/QARequHealing.htm
Pennsylvania guidelines and links to guidelines developed by appropriate professional organizations
Fact sheets on:
  “Minimum QA Requirements for Healing Arts Radiography”
  “Model QA Guidelines for Dental, Diagnostic Radiology and Mammography”

## Slide 77: Regulations
Part 5

## Slide 78: Brief History and Overview of Radiation Regulations
X-ray machines regulated by states since their early development.
Shortly after WW II, most radioactive materials and all nuclear reactors assigned to be regulated by federal government.
  Originally under the Atomic Energy Commission (AEC) and later under the Nuclear Regulatory Commission (NRC).

## Slide 79: Agreement State Status
Because of overlap of many regulatory issues regarding X-rays and radioactive materials, Agreement State arrangements were made.
Federal agency (AEC and later NRC) would allow state regulatory oversight of most radioactive materials.
Federal government maintained control over reactors and certain materials dealing with defense.

## Slide 80: Pa. Agreement State Status
Pennsylvania became Agreement State in 2008.
State regulations found in 25 Pa. Code – Environmental Protection, Chapters 215-240.
Federal regulations from Title 10 of the Code of Federal Regulations (or more commonly: 10 CFR) incorporated into some Pennsylvania regulations by reference.

## Slide 81: Role of FDA
Federal Food and Drug Administration (FDA) responsible for protecting public from hazardous or unnecessary exposure to radiation-emitting electronic products.
Oversees manufacturer compliance and studies biological effects of radiation.
FDA regulations found in 21 CFR.

## Slide 82: FDA, DEP and the MQSA
FDA also regulates mammography facilities under the Federal Mammography Quality Standards Act (MQSA).
In Pennsylvania, DEP’s Bureau of Radiation Protection contracts with FDA to perform annual inspections of mammography facilities.
Pennsylvania also maintains list of certified mammography facilities.

## Slide 83: Regulations Related to Use of Diagnostic X-rays
Following slides briefly review regulations regarding diagnostic X-rays in 25 Pa. Code and the incorporated section of 10 CFR.
Not intended to be comprehensive; registrants should always refer directly to the regulations for issues at their facilities.

## Slide 84: Registration of Radiation-Producing Machines
Provides requirements for registration, renewal, expiration, or termination of certificate of registration and transfer or disposal of machine.
Applications for renewal sent out at least 2 months prior to expiration.
DEP must be notified of transfer or disposal of X-ray devices.
25 Pa. Code Chapter 216

## Slide 85: Standards for Protection Against Radiation
Largely parallels federal regulations in 10 CFR Part 20, which are incorporated by reference.
Many sections not relevant to X-ray facilities since they apply to radioactive materials.
Will highlight those that apply to medical X-ray devices.
25 Pa. Code Chapter 219

## Slide 86: Occupational Dose Limits
Applies only to occupational workers – individuals exposed in course of their work.
Does not include background radiation, medical administration to the worker, or other exposures as a member of the public.
Note: this means that radiation doses to an individual from medical procedures performed on them do not fall under these regulations – a frequent concern of radiation workers.
25 Pa. Code §§ 219.21-22
Occupational Dose Limits

## Slide 87: Annual Occupational Dose Limits for Diagnostic X-ray Workers with no Other Occupational Exposure
Effective Dose Equivalent to Whole Body | 5 rem (0.05 Sv)
Lens of Eye Dose Equivalent | 15 rem (0.15 Sv)
Shallow dose equivalent to skin of whole body or extremity | 50 rem (0.5 Sv)
Dose limits for minors (occupational) (< 18 years old) | 10 % of adult limits
Dose equivalent to embryo/fetus of declared pregnant female (See NRC Reg. Guide 8.13) | 0.5 rem (5 mSv) during entire pregnancy (See 10CFR20.1208 for additional information and guidelines)

## Slide 88: Dose Limits for Public
Shall not exceed 0.1 rem (1 mSv) in a year.
Excludes background radiation or medically administered radiation.
Example: This limit applies to a patient (or other family member) in a waiting room, but not to the medical treatment of the patient.
Note: Value was changed from 0.5 rem in 1990s; Pennsylvania does not require retrofitting of shielding for installations existing before Nov. 18, 1995 as long as similar equipment is used.
25 Pa. Code § 219.51

## Slide 89: Shielding of X-ray Facilities
X-ray facilities are shielded to protect individuals in adjoining areas and are based on expected use of those areas outside the X-ray room.
Major renovations of facilities or replacement of machines resulting in higher workloads may require a review of the shielding in-place to ensure that it is still adequate.

## Slide 90
Sources of radiation (including X-ray machines) shall be secured from unauthorized removal or access while in storage or available for use.
Storage and Control
25 Pa. Code §§ 219.131-132

## Slide 91
Radiation-producing machines are required to be labeled indicating that radiation is produced when energized:
CAUTION – RADIATION
THIS EQUIPMENT PRODUCES RADIATION WHEN ENERGIZED
(Caution signs based on radiation level not required for rooms with machines used solely for diagnosis in the healing arts).
Posting Requirements
25 Pa. Code §§ 219.159-160

## Slide 92
Report shall be made to the state of stolen, lost or missing sources of radiation including radiation-producing machines.
Report required to state if determination by physician of actual or suspected damage to organ or system of patient exposed to therapeutic or diagnostic radiation.
Reporting Requirements
25 Pa. Code §§ 219.221,222,229

## Slide 93: Notices, Instructions & Reports to Workers; Inspections &Investigations
Registrants required to post:
  Pa. Code Chapters 219 and 220
  Certificate of Registration
  Applicable operating procedures
  Notices of violations
Alternatively, registrant may post notice describing documents and where they may be examined.
25 PA Code Chapter 220
25 PA Code Chapter 220

## Slide 94
DEP Form 2900-FM-RP0003 “Notice to Employees” required to be posted.
Available on DEP’s website.
Outlines employer’s and worker’s responsibilities, items covered by regulations, reports on workers’ radiation history and inspections.
Provides contact information for DEP’s Bureau of Radiation Protection.
Notices to Employees”

## Slide 95
2900-FM-RP0003 Rev. 8/2008
COMMONWEALTH OF PENNSYLVANIA
DEPARTMENT OF ENVIRONMENTAL PROTECTION
BUREAU OF RADIATION PROTECTION
NOTICE TO EMPLOYEES
STANDARDS FOR PROTECTION AGAINST RADIATION; NOTICES, INSTRUCTIONS
AND REPORTS TO WORKERS; INSPECTIONS; EMPLOYEE PROTECTION
In Title 25 of its Rules and Regulations, the Pennsylvania Department of Environmental Protection has established standards for your protection against radiation hazards and has established certain provisions for the options of workers engaged in work under a Department license or registration.
YOUR EMPLOYER’S RESPONSIBILITY
Your employer is required to:
Apply these Department of Environmental Protection regulations and any conditions of your employer’s radioactive materials license to all work involving radiation sources.
Post or otherwise make available to you a copy of the Department of Environmental Protection regulations, licenses, and operating procedures which apply to work in which you are engaged, and explain their provisions to you.
Post Notice of Violation involving radiological working conditions, proposed imposition of civil penalties and orders.
YOUR RESPONSIBILITY AS A WORKER
You should familiarize yourself with these provisions of the Department of Environmental Protection regulations and operating procedures which apply to the work in which you are engaged. You should observe their provisions for your own protection and protection of your co-workers. If you observe a violation or possible safety concern, you should report it immediately to your supervisor or contact DEP. You may be personally subject to enforcement action if through deliberate misconduct you cause or attempt to cause a violation of DEP requirements or deliberately provide inaccurate or incomplete safety information to DEP or your employer.
WHAT IS COVERED BY THESE REGULATIONS
Limits on exposure to radiation and radioactive materials in restricted and unrestricted areas.
Measures to be taken after accidental exposure.
Personal monitoring, surveys, and equipment.
Caution signs, labels, and safety interlock equipment.
Exposure records and reports.
Options for workers regarding Department inspections.
Related matters.
REPORTS ON YOUR RADIATION HISTORY
1. The Department of Environmental Protection regulations require that your employer give you a written report if you receive an exposure in excess of any applicable limit as set forth in the regulations or the license. The basic limits for exposure to employees are set forth in Chapter 219 of the regulations. This chapter specifies limits on exposure to radiation and exposure to concentrations of radioactive material in air.
2. If you work where personal monitoring is required pursuant to Chapter 219:
(a) Your employer must advise you annually of your exposure to radiation, and
(b) You may request a written report of your radiation exposure when you leave your job.

## Slide 96
INSPECTIONS
All activities involving radiation are subject to inspection by representatives of the Pennsylvania Department of Environmental Protection. In addition, any worker or representative of workers who believes that there is a violation of the Department regulations of the terms of the employer’s license or registration with regard to radiological working conditions in which the worker is engaged, may request an inspection by sending a notice of the alleged violation to the Bureau of Radiation Protection. The request must set forth the specific grounds for the notice, and must be signed by the worker as the representative of the workers or their self. During inspections, Department inspectors may confer privately with workers, and any worker may bring to the attention of the inspectors any past or present condition which that worker believes contributed to or caused any violation as described above.
INQUIRIES
Inquiries dealing with matters outlined above or other reports and correspondence can be sent to the Bureau of Radiation Protection, Pennsylvania Department of Environmental Protection, P.O. Box 8469, Harrisburg PA 17105-8469.
Telephone (717) 787-3720
Facsimile (717) 783-8965
Off hours emergency call PEMA: (717) 651-2001

## Slide 97: X-rays in the Healing Arts
Chapter 221 provides detailed information on requirements for medical X-rays
  General Provisions (§§ 221.1 and 221.2) discuss purpose and scope of this chapter and provide extensive list of definitions.
  Administrative Controls (§§ 221.11-221.15) cover registrant responsibilities, reports, records, applicability and associated information. These controls have largely been covered during the section on X-ray machines.

## Slide 98: X-rays in the Healing Arts
  Diagnostic Installations General Requirements (§§ 221.21-221.49) provide specific regulations for operation, maintenance, and control of X-rays used in healing arts.
  Registrant should review these regulations for their specific facility to determine application.
  A table of the section titles and a brief review of the requirements follows:

## Slide 99
Section Number and Title | Abbreviated summary of section requirements (See Regulations for Additional Details)
221.21 Diagnostic equipment requirements | Certified components shall comply with relevant regulations of the Food and Drug Agency (21CFR 1020.30 – 1020.33)
221.22 Battery charge indicator | Control panels on battery powered x-ray generators shall visually indicate proper battery operation
221.23 Leakage radiation from diagnostic source assembly | May not exceed 100 mR in one hour at 1 meter
221.24 Radiation from components other than diagnostic source assembly | May not exceed 2 mR in 1 hour at 5 cm from accessible surface
221.25 Beam Quality | Table I gives minimum filtration requirements based on operating voltage. Table II gives minimum HVL values that will meet these requirements
221.26 Multiple tubes | When multiple tubes are controlled by one switch, indicators on the control panel and at or near the tube housing assembly shall indicate which tube has been selected

## Slide 100
221.27 Mechanical support of tube head | Tube housing assembly shall remain stable during exposure (unless movement is a designed function of system)
221.28 Technique indicators | Technique factors shall be indicated (except for automatic exposure controls in which case mAs shall be indicated). Equipment having fixed technique factors may indicate them with permanent marking on equipment
221.29 Kilovoltage (kV) accuracy | Output for variable kV units may not vary from set-indicated value by more than 10% Output for fixed kV units may not vary from set-indicated value by more than 20%
221.30 Exposure reproducibility for noncertified systems | Coefficient of variation of exposure reproducibility may not exceed 0.1 when technique factors held constant. (See definitions in 221.2 for formula for this calculation)
221.31a Locks | Position locking, holding and centering devices shall function as intended
221.32a Radiographic beam limitations | Useful beam shall be limited to area of clinical interest. Specifics are given for beam limiting devices regarding accuracy, adjustment, and alignment. Intraoral dental system requirements for beam limitation are specified
221.33a Radiation from capacitor energy storage equipment in standby status | When switch or timer not activated, may not exceed 2 mR/hour at 5 cm from accessible surface when fully charged and beam limiting device fully open

## Slide 101
221.34a Radiation exposure control | Requirements to ensure exposure controls are given including switch operations, visible and audible signals and other requirements for manual and automatic exposure control. Also stationary systems shall have controls in protected area and require operator to remain there; mobile and portable units shall be designed so operator is at least 2 meters from patient and x-ray tube head when operating system.
221.35a Fluoroscopic x-ray systems | Fluoro- systems shall use image intensifier and comply with prior regulations in this chapter.
221.36a Limitation of useful beam of fluoroscopic equipment | Requirements are given for primary protective barrier placement, adjustment and size of the x-ray field, minimum source to skin distance, and spot image device requirements
221.37a Activation of fluoroscopic tube | Dead-man switch and means to terminate serial images shall be provided

## Slide 102
221.38a Entrance Exposure Rate | Entrance exposure rates, frequency of measurements and compliance requirements are given Entrance exposure rates are: 10 R/min. for systems without high level control 20 R/min. for systems with high level control activated 10 R/min. for systems with high level control, but not activated
221.39a Barrier transmitted radiation rate limits | Protective barrier may not transmit >2mR/hr at 10 cm from accessible surface of fluoroscopic imaging assembly for each R/min. of entrance exposure rate
221.40a Indication of tube voltage and current | During fluoroscopy and cinefluorography, voltage and current shall be indicated
221.41a Fluoroscopic timer | Timing device activated by fluoroscopic switch shall be provided. It shall provide audible signal or temporary/permanent interruption when preset limit not exceeding 5 minutes is reached
221.42a Control of scattered radiation | Limits for scatter radiation originating either under or above the table top are specified
221.43a Mobile fluoroscopes | In addition to other fluoroscopic requirements, shall provide image intensification

## Slide 103: X-rays in the Healing Arts
Remainder of Chapter 221 deals with radiation therapy simulations systems, therapeutic X-rays systems with energies less than 1 MeV and computed tomography X-ray systems and are beyond the scope of this training program.

## Slide 104: Conclusion
This section has presented an overview of the regulations that apply to medical X-rays used in low-risk procedures. Registrants should directly consult the applicable regulations or contact the Pennsylvania Department of Environmental Protection, Bureau of Radiation Protection, for questions or concerns.

## Slide 105: Online Continuing Education Radiation Safety Quiz
The Online Continuing Education Radiation Safety Quiz is found at
http://www.dep.pa.gov/Business/RadiationProtection/RadiationControl/X-rayMachineProgram/Pages/Continuing-Education-In-Radiation-Safety-Quiz.aspx
Upon successful completion of the quiz, print out the confirmation webpage and retain in your records.
25 PA Code Chapter 220

## Slide 106: Contact Information
Pa. Department of Environmental Protection
Bureau of Radiation Protection
P.O. Box 8469
Harrisburg, PA 17105-8469
Telephone (717) 787-3720
Email: RA-EPRadiationContro@pa.gov
Off-hours emergency call PEMA (717) 651-2001
Notes: Took out Fax, and added email.